# Software Verification Report — Monitoring Waveform Acceptance Check

## 1. Document Control

| Field | Value |
| --- | --- |
| Software | Monitoring Waveform Acceptance Check |
| Company | Amber Therapeutics |
| Version reviewed | 1.0 (APP_VERSION in signal_viewer.py, FileVersion 1.0.0.0 in version_info.txt) |
| File(s) reviewed | signal_viewer.py (the application) |
| File(s) out of scope | plot_batch.py (a separate, unvalidated developer plotting utility — not part of the acceptance-check application, shares no code with it) |
| Review method | Static code review / requirements traceability. No automated tests were executed against this version — see §5. |
| Reviewed by | Vinil T Chackochan (viniltc@gmail.com), with AI-assisted review (Claude) |
| Date | 2026-08-28 |

## 2. Purpose and Scope

This report verifies, by inspection of the source code, that signal_viewer.py correctly implements its stated acceptance criteria and data-integrity rules, and behaves safely on invalid input. It is intended as audit evidence that the logic governing PASS/FAIL verdicts matches the documented acceptance limits, and that failure paths do not silently corrupt or misreport results.

This report does not constitute dynamic/execution-based testing (no synthetic waveforms were run through the tool). It should be read as a design/code review, and paired with executed test evidence (Approach 1 — an automated test suite) before being relied on as complete verification. See §7.

## 3. Requirements Traceability Matrix

Requirements below are derived from the constants, docstrings, and in-code comments in signal_viewer.py itself — there is no separate external requirements document. Each is checked against its implementation.

### 3.1 Acceptance criteria (PASS/FAIL logic)

| ID | Requirement | Implementation | Review verdict |
| --- | --- | --- | --- |
| R1 | Accepted waveform shape is triangle only | LIMIT_WAVEFORM = "triangle" (L66); shape identified from RMS/peak ratio against WAVEFORM_RMS_RATIO (L76-L80), compared in evaluate() (L176-L182) | Verified. Ratio constants (1/√3 triangle, 1/√2 sine, 1.0 square) match standard waveform RMS formulas. |
| R2 | Frequency target 10.0 Hz ± 0.5 Hz | LIMIT_FREQUENCY_HZ/LIMIT_FREQUENCY_TOLERANCE_HZ (L67, L71); range check in evaluate() (L185-L194) | Verified. Both above- and below-range and “undetermined” (NaN) cases produce a FAIL with a specific reason. |
| R3 | Peak-to-peak amplitude target 10.0 mV ± 1.0 mV | LIMIT_PEAK_TO_PEAK_MV/LIMIT_AMPLITUDE_TOLERANCE_MV (L68, L70); range check (L196-L204) | Verified. |
| R4 | No single cycle period may deviate more than 1.0 ms from the nominal 100 ms period | LIMIT_TIME_TOLERANCE_MS (L69); worst-cycle deviation computed in _period_deviation() (L276-L291) against every individual cycle (not just an average), checked in evaluate() (L206-L214) | Verified. Uses interpolated zero-crossing times (L294-L317), so resolution isn’t limited to the sample interval. |
| R5 | A PASS requires all four checks (shape, frequency, amplitude, timing) to succeed; any failing check must state why | evaluate() accumulates a reasons list from all four checks before deciding; only returns PASS if the list is empty (L216-L218) | Verified. Failing on multiple axes reports all reasons, not just the first. |
| R6 | Acceptance limits must not be editable from the running application; changing them requires a rebuild and version bump | Limits are module-level constants, displayed read-only via ttk.Label (not Entry) in the UI panel (L679-L692); no code path writes to the LIMIT_* constants at runtime | Verified. |
| R7 | The application must fail visibly at startup if the fixed limits are internally inconsistent, rather than producing wrong verdicts silently | configured_limits() validates the waveform name and that all tolerances are positive numbers, raising ValueError (L127-L157); called once in main() before the window is created (L1035) and caught by the top-level handler (L1044-L1067) | Verified. |

### 3.2 File format and data-integrity checks

| ID | Requirement | Implementation | Review verdict |
| --- | --- | --- | --- |
| R8 | A batch must contain exactly 4 CSV files | EXPECTED_FILES = 4 (L52); checked before any file is parsed (L912-L924) | Verified. Wrong count blocks the check entirely and lists what was found. |
| R9 | Each file must have 4 metadata lines followed by a column header naming Time then Signal | read_recording() reads N_META_LINES + 1 lines, rejects short files (L366-L370), and validates the header text (L376-L382) | Verified. |
| R10 | Time values must strictly increase (no corrupt/reordered rows) | if np.any(np.diff(time) <= 0): raise FormatError (L422-L426) | Verified. |
| R11 | The declared sample rate (metadata line 4) should match the actual time step; a mismatch >1% is flagged | Cross-check at L429-L435 | Verified, with one minor observation — see §6.1. |
| R12 | Non-numeric sample rows are dropped, not silently included, and the operator is told how many | data.isna().any(axis=1).sum() counted before dropna(), warning appended if any (L406-L419) | Verified. |
| R13 | A file that cannot be opened, decoded, or parsed produces a specific, named failure rather than crashing the app | FormatError raised with a distinct message for each failure mode (OSError, UnicodeDecodeError, CSV parse Exception) (L358-L364, L403-L404); caught per-file in run_check() (L928-L937) | Verified. |

### 3.3 Batch-level consistency (non-blocking warnings)

| ID | Requirement | Implementation | Review verdict |
| --- | --- | --- | --- |
| R14 | All files in a batch should be from the same subject | validate_batch() (L445-L449) | Verified. Warning only, does not block the check — reasonable, since subject mismatch doesn’t invalidate individual channel measurements. |
| R15 | Channel labels should be unique within a batch | L451-L454 | Verified. |
| R16 | Recording durations across the batch should agree within 1% | L456-L462 | Verified. |

### 3.4 Result integrity and reporting

| ID | Requirement | Implementation | Review verdict |
| --- | --- | --- | --- |
| R17 | Results on screen (plots, table, verdict banner) must be cleared before every new check, so a failed run can never leave stale data next to a new error | _clear_results() called unconditionally at the top of run_check() (L909), and again from choose_folder() | Verified. |
| R18 | A PDF report is produced automatically for every completed check, with a unique filename | _save_report() → next_report_path() increments a numeric suffix on collision (L592-L600) | Verified. |
| R19 | If the PDF cannot be written (locked file, folder permissions), the already-displayed results must remain valid and the user must be warned, not blocked | _save_report() is called after plot()/_fill_table()/_set_verdict_banner() have already run (L963-L967); write failures are caught (PermissionError, generic Exception) and shown as a non-blocking messagebox.showwarning (L834-L848) | Verified. |
| R20 | The report must carry software identity and the run timestamp for traceability | PDF metadata (Title, Creator) set from APP_TITLE/APP_VERSION (L509-L513); summary page header includes run time, subject, recorded time, software version (L523-L530) | Verified. |
| R21 | Measurements shown to the operator are computed on the full recording, never on a truncated/windowed view, even when the plot itself is windowed for readability | measure() is called on the untruncated rec.data (L953-L954); only plot() truncates to window_ms for display (L998-L1000); the PDF and on-screen table both carry an explicit caption stating the plotted span is not the measured span (L496-L499, L765-L771) | Verified. |

## 4. Code Review Checklist (general robustness)

| Check | Result |
| --- | --- |
| Application-level exceptions during startup are caught and shown to the user (no silent crash in a windowed build with no console) | Pass — L1044-L1067 |
| No bare except: clauses that could mask unexpected errors | Pass — all except clauses name specific exception types or explicitly re-raise/report via exc |
| No mutable global state used for the acceptance limits (only read via configured_limits()/constants) | Pass |
| File paths and folder creation handle OSError explicitly rather than assuming success | Pass — L818-L826 |
| No use of eval, exec, shell invocation, or other injection-prone constructs | Pass |
| Version number (APP_VERSION) is co-located with the fixed limits with an explicit comment to keep them in step | Pass — L38-L64 |

## 5. What this review does not cover

- No execution/dynamic testing. No synthetic waveform (known-good, known-bad, boundary-value) was actually run through measure()/evaluate() to confirm the numeric output is correct — this review confirms the logic matches the stated limits, not that the arithmetic is bug-free under real data (e.g., correctness of the FFT fallback frequency estimate, or behavior with irregular sample spacing).

- No UI/usability testing (see the separate architecture discussion earlier in this session for how the UI is built).

- No review of the PyInstaller packaging (build.bat, MonitoringWaveformAcceptanceCheck.spec) — packaging integrity is out of scope for this pass.

## 6. Observations

### 6.1 Minor — sample-rate cross-check skipped when declared rate is exactly 0

L429 uses if declared_rate:, which is falsy for 0.0. A file that declares a sample rate of exactly 0 Hz (rather than an unparsable value) would skip the cross-check silently instead of generating a warning. Very low likelihood in practice (a real recorder would not declare 0 Hz), and it fails safe (skips a warning, not a required check), so no action is required unless the team wants stricter handling.

### 6.2 Recommendation — pair this report with executed tests

This is a static/code-review verification. It is strong evidence the implementation matches its stated criteria, but per §5 it is not a substitute for running known-input test vectors through the tool and recording actual output. See §7 for an executed run that partially closes this gap; a synthetic-input pytest suite (ideal signal, boundary-value signals just inside/outside each limit, malformed files) is still recommended for full arithmetic-level coverage of measure()/evaluate()/read_recording().

## 7. Executed Test Evidence (Dynamic Check)

In addition to the static review above, the acceptance-check logic was executed against a set of real recorded batches using tools/run_batch_checks.py — a runner that imports and calls the exact same functions used by the application (read_recording, validate_batch, measure, evaluate), so the result is evidence about the application’s actual behaviour, not a reimplementation.

| Field | Value |
| --- | --- |
| Date run | 2026-08-28 |
| Batches checked | 9 (one subfolder per batch, each with the expected 4 CSV files) |
| Result | 3 PASS, 6 FAIL, 0 SKIPPED/ERROR |
| Raw per-batch and per-channel output | results/batch_check_summary.csv, results/batch_check_channels.csv (not committed — derived from identifiable subject data; see .gitignore) |

Findings from this run: - Measured peak-to-peak amplitude scaled monotonically and consistently with the source recordings’ stated stimulus amplitude across all batches, and measured frequency tracked the stated target frequency (including correctly failing an off-target ~9 Hz batch). This is evidence the frequency and amplitude measurement logic (§3.1, R2–R3) produces sensible, consistent results on real hardware output, not just on the constants it was reviewed against. - Two batches contained individual channels with a large single-cycle timing deviation (tens of ms, against sibling channels in the same batch measuring a normal ~0.14 ms) — correctly flagged as FAIL by R4/R16. Because the deviation appears on isolated channels within an otherwise-clean batch (same code, same run), this looks like a genuine recording-level artifact rather than a defect in the timing-deviation logic itself, but the underlying recordings should be reviewed by someone familiar with the acquisition hardware to confirm the cause. - No file in this run triggered FormatError or was skipped for the wrong file count, so this run does not add coverage of the file-format failure paths (R9–R13) — those remain verified by static review only (§3.2).

## 8. Conclusion

Based on static code review, all identified acceptance-criteria and data-integrity requirements (R1–R21) in signal_viewer.py v1.0 are correctly implemented as stated in the code’s own constants and comments. One low-risk, fail-safe observation was noted (§6.1); no defects were found that would cause an incorrect PASS/FAIL verdict. An executed run against real data (§7) corroborates the amplitude and frequency measurement logic and demonstrates the timing check correctly detects a real-world anomaly; file-format failure paths remain verified by static review only.

| Role | Name | Signature | Date |
| --- | --- | --- | --- |
| Reviewer |  |  |  |
| Approver |  |  |  |